| TECPEL CO., LTD. Model: HP-100 |  |
| --- | --- |
| Bandwidth | DC to 100MHz (-3dB) |
| Attenuation | 1:100/1000 |
| Accuracy | +/- 2% |
| Rise Time | 3.5ns |
| Input Impedance | 4Mohm//7pF each side to ground |
| Input Voltage |  |
| -Max. Differential* | ±140V(DC + Peak AC) @ 1/100 ±1400V (DC + Peak AC) @ 1/1000 |
| - Common Mode Range* | ±1400V (DC + Peak AC) @ 1/100 & 1/1000 |
| - Absolute Max. Voltage* (Differential or Common Mode) | ±1400V (DC + Peak AC) @ 1/100 & 1/1000 |
| Output Voltage |  |
| - Swing (into 50kohm load) | ±1.4V |
| -Offset (typical) | <+/-5mV |
| -Noise (typical) | 0.6mVrms |
| -Source Impedance (typical) | 50 ohm |
| CMRR (typical) | -80dB@60Hz; -50db@1MHz |
| Ambient Operating Temperature | -10 to 40 degree centigrade |
| Ambient Storage Temperature | -30 to 70 degree centigrade |
| Ambient Operating Humidity | 25 to 85% RH |
| Ambient Storage Humidity | 25 to 85% RH |
| Power Requirements |  |
| -Standard | 4xAA cells or 6VDC/200mA mains adaptor** or regualted 9VDC/120mA mains adaptor |
| -Options | Power leads |
| Length of Input Leads | 30cm |
| Length of BNC Cable | 90cm |
| Weight | 500gms |
| Dimension (LxWxH) | 207mmx83mmx38mm |

* Voltage limit is the lesser of the DC+Peak AC and RMS values.. ** a. The supplied voltage must be less than 12V and greater than 4.4V, otherwise the probe could be damaged or can't be operated properly. b. polarity is “+” inside and “–” outside. For wrong polarity, built-in circuit protects the probe, no danger or damage will occur. c. When the voltage of the cells become too low, the power indicator on the panel will flicker.

| TECPEL CO., LTD. Model: HP-025 |  |
| --- | --- |
| Bandwidth | DC to 25MHz (-3dB) |
| Attenuation Ratio | 1:20/200 |
| Accuracy | +/- 2% |
| Rise Time | 14ns |
| Input Impedance | 4Mohm//5.5pF each side ground |
| Input Voltage |  |
| -Differential Range* | +/-140V(DC+Peak AC) or 140Vrms @1/20 +/-1400V(DC+Peak AC) or 1000Vrms @1/200 |
| - Common Mode Range* | +/-1400V(DC+Peak AC) or 1000Vrms @1/20 & 1/200 |
| - Absolute Max. Voltage* (Differential or Common Mode | +/-1400V(DC+Peak AC) or 1000Vrms CAT III @1/20 & 1/200 |
| Output |  |
| - Swing (into 2k ohm load) | +/-7V |
| -Offset (typical) | <+/-5mV |
| -Noise (typical) | 0.7mVrms |
| -Source Impedance (typical) | 1 ohm @1kHz, 8 ohm @1MHz |
| CMRR (typical) | -80dB@50Hz,-60dB@20kHz |
| Ambient Operating Temperature | -10 to 40 degree centigrade |
| Ambient Storage Temperature | -30 to 70 degree centigrade |
| Ambient Operating Humidity | 25 to 85% RH |
| Ambient Storage Humidity | 25 to 85% RH |
| Power Requirements |  |
| -Standard | 4xAA cells or 6VDC/60mA mains adaptor** or regualted 9VDC/40mA mains adaptor** |
| -Options | Power leads |
| Length of BNC Cable | 95cm |
| Length of Input Leads | 45cm |
| Weight | 400gms (probe and PVC jacket) |
| Dimensions(LxWxH) | 170mm x 63mm x 21mm |

* Voltage limit is the lesser of the DC+Peak AC and RMS values.. ** a. The supplied voltage must be less than 12V and greater than 4.4V, otherwise the probe could be damaged or can’t be operated properly. b. polarity is “+” inside and “–” outside. For wrong polarity, built-in circuit protects the probe, no danger or damage will occur. c. When the voltage of the cells become too low, the power indicator on the panel will flicker.

TECPEL CO., LTD.

No. 13, Alley 5, Lane 31, Chung_Teh Street,

Taipei, Taiwan, R. O. C.

Tel: 886-2-2737-5866 ; Fax: 886-2-2737-3343

E-mail: sales@tecpel.com

Web: http://www.oka.com.tw  ;  http://www.tecpel.com/